## Small Molecules & Polymers
Technique | Abbreviation | Sample Type | Main Use | Sample Requirements | Sample Requirements | Sample Requirements | Sample Requirements | Sample Requirements | Sample Requirements | Instrument Parameters | Typical Application | Important
|  |  |  | Amount (mg) | Physical State | Solubility | Sample Preparation | Purity | Sensitivity | Detector
High Performance Liquid Chromatography | HPLC | Small molecules Polymers | Separation quantification | 0.1-10 mg | Solution in appropriate solvent | Water Acetonitrile Methanol Isopropanol n-Propanol Ethyl acetate Dichloromethane Hexane Cyclohexane | 1. Remove major contaminant groups such as proteins and salts 2. 1-2 microL sample | Free of particles filtered | ng-μg | UV-Vis RID | Pharmaceuticals metabolites | 1. UV-Vis absorption for UV detection 4. Small molecules - Reverse phase HPLC 2. Polar molecules - Hydrophilic interaction chromatography (HILIC) or normal phase 3. Ionic compounds - Ion exchange chromatography 4. Derivatization needed for some molecules
Ultra-High Performance Liquid Chromatography | UHPLC | Small molecules Polymers | Separation quantification | 0.01-1 mg | Solution in appropriate solvent | Water Acetonitrile Methanol Isopropanol n-Propanol Ethyl acetate Dichloromethane Hexane Cyclohexane | 1. Remove major contaminant groups such as proteins and salts 2. <1 microL sample | Free of particles filtered | pg-ng | UV-Vis RID | Pharmaceuticals metabolites | High resolution and faster analysis compared to HPLC; suitable for complex mixtures
Liquid Chromatography | LC | Small molecules Polymers | Separation quantification | 0.1-10 mg | Solution in appropriate solvent | Water Acetonitrile Methanol Isopropanol n-Propanol Ethyl acetate Dichloromethane Hexane Cyclohexane | Remove particulate matter | Free of particles | ng-μg | UV-Vis RID | Pharmaceuticals metabolites | Broad applicability; includes various modes like reverse-phase normal-phase and ion-exchange
Two-Dimensional Liquid Chromatography | 2D-LC | Small molecules Polymers | Separation quantification | Tens of microliters | Solution in appropriate solvent | Solvents in both dimensions must be compatible and miscible | Remove particulate matter | Free of particles | pg-ng | UV-Vis MS RID | Complex samples proteomics metabolomics | 1. Combines two orthogonal separations 2. High resolution for complex mixtures
Ion Chromatography | IC | Small molecules Polymers | Separation quantification of anions and cations | ~0.01–1 mg or mL | Aqueous solution | Water-soluble | Remove particulates filter adjust pH if necessary | Free of particles | ng–µg | Conductivity UV or MS | Environmental analysis water quality pharmaceuticals | 1. Ensure low background conductivity 2. Suppressor may be required for IC-MS
Interaction Chromatography (Hydrophilic) | HILIC | Small polar molecules biomolecules | Separation of polar/ionic compounds (not retained well in reversed-phase) | ~0.01–1 mg | Aqueous-organic solution | Water-miscible solvents (e.g. acetonitrile methanol) | Filtered diluted in high organic content | Free of particles | ng–µg | UV-Vis MS | Metabolomics pharmaceutical analysis glycan profiling | High organic content mobile phase enhances retention of polar analytes
Gas Chromatography | GC | Small molecules | Separation of volatiles | 0.1-5 mg | Volatile heat-stable (boiling point <400°C) thoroughly dry | Hexane Dichloromethane Ethyl acetate Methanol | Volatile ionizable | High purity | ng-μg | Flame Ionization Detector (FID) Thermal Conductivity Detector (TCD) MS | Environmental forensics | Suitable for volatile and semi-volatile compounds; requires thermal stability
Thin Layer Chromatography | TLC | Small molecules | Qualitative separation | ~0.001 mg | Dissolved in volatile solvent | In eluting solvents | Spot on TLC plate with capillary or micropipette | Sufficient for visual detection | μg | Visual UV fluorescence | Purity checks | Simple cost-effective method for quick qualitative analysis
Capillary Electrophoresis | CE | Small molecules charged biomolecules | Separation of ions | 0.01-1 mg | Aqueous solution | Aqueous buffer system | Filtered pH adjusted buffer solutions | Free of particles | ng | UV-Vis Conductivity | Amino acids peptides | Sensitive to electroosmotic flow and pH excellent for ionic species
Nuclear Magnetic Resonance Spectroscopy | NMR | Small molecules Polymers | Structure elucidation quantification | 1-10 mg | Liquid (soluble in deuterium-based solvents) | Deuterated chloroform (CDCl₃) Deuterated DMSO (DMSO-d₆) Deuterated methanol (CD₃OD) Deuterated water (D₂O) Deuterated acetone (acetone-d₆) | Dissolve in deuterated solvent filter if necessary | >95% free from paramagnetic impurities iron filings and dust | μmol | NMR detector | Research QC | 1. Requires deuterated solvents 2. Non-destructive technique with detailed structural output
Mass Spectrometry | MS | Small molecules Polymers | Identification quantification structure | 0.01-1 mg | Depends on ionization source (solid liquid or solution) | Depends on ionization method (e.g. ESI: polar solvents; MALDI: matrix-assisted) | Method-specific: matrix for MALDI solvent for ESI/APCI filtration if needed | High purity | Sensitivity depends on instrument configuration and ionization method typically pg-ng | TOF Orbitrap Quadrupole | Metabolomics drug analysis | 1. Requires ionizable analysis 2. Very sensitive 3. Ideal for exact mass and structure info
Liquid Chromatography Mass Spectrometry | LC-MS | Small molecules Polymers | Identification quantification | 0.01-1 mg | Completely dissolved solution | Water Acetonitrile Methanol Isopropanol n-Propanol Ethyl acetate Dichloromethane Hexane Cyclohexane | Remove particles adjust solvent to be compatible with ESI or APCI source | Free from mechanical impurities | Sensitivity depends on instrument configuration and ionization method typically pg-ng | UV + MS detector | Metabolomics drug analysis | 1. Combines LC separation with high sensitivity MS detection 2. Ideal for complex sample matrices
Gas Chromatography Mass Spectrometry | GC-MS | Small molecules | Identification quantification | 0.01-1 mg | Vaporizable within column temperature range | Hexane Dichloromethane Ethyl acetate Methanol |  | High purity | Sensitivity depends on instrument configuration and ionization method typically pg-ng | MS (Electron impact or chemical ionization) | Environmental forensics | 1. Excellent for volatile organic 2. Provides both separation and structure-specific detection
Tandem Mass Spectrometry | MS/MS | small molecules Polymers | Structural elucidation | 0.001-0.1 mg | Depends on ionization source | Polar solvents: water methanol acetonitrile depending on source | Filter dilute to analytical concentration remove salts if needed | High purity | Sensitivity depends on instrument configuration and ionization method typically fg-pg | Two or more mass analyzers in series (e.g. QqQ QTOF ion trap) | Metabolomics drug analysis compound identification | Enables fragmentation of selected ions for structural info; ideal for identifying unknowns quantifying knowns and confirming molecular structure; can be coupled with LC or GC (LC-MS/MS GC-MS/MS)
High-Resolution Mass Spectrometry | HRMS | Small molecules biomolecules polymers | Accurate mass measurement | 0.001-0.1 mg | Depends on ionization source (solid liquid or solution) | Polar solvents (e.g. water methanol acetonitrile) or matrix for MALDI | Dissolve in suitable solvent filter remove salts | High purity | Sensitivity depends on instrument configuration and ionization method typically fg-pg | TOF Orbitrap FT-ICR | Research metabolomics polymer characterization | High mass accuracy and resolution for complex mixtures and exact mass determination
Ultraviolet-Visible Spectroscopy | UV-Vis | Small molecules Polymers | Detection quantification | 0.1-10 mg | Solution or transparent solid | Water Methanol Ethanol Acetonitrile Chloroform Toluene | Dilute to appropriate absorbance (A = 0.1–1.0); remove bubbles and particulates | Optical clarity important | μg-mg | Photodiode array or single-beam photodetector | Purity ID | Fast simple and non-destructive; Beer's Law relates absorbance to concentration; baseline correction critical for accurate results
Fourier Transform Infrared Spectroscopy | FTIR | Small molecules Polymers | Functional group analysis | 0.1-10 mg | Solid liquid or solution; thoroughly dried | Water Methanol Ethanol Acetonitrile Chloroform Toluene | Solid: KBr pellet ATR or film on substrate; liquid: sandwiched between IR-transparent plates | Pre-purified free from moisture | μg-mg | MCT (Mercury Cadmium Telluride) or DTGS detector | Purity ID | Fast non-destructive; ATR (attenuated total reflection) is widely used for ease of sample prep; good complement to NMR and MS for structural elucidation
Microscale Infrared Spectroscopy | Micro-IR | Small molecules Polymers | Identification and spatially-resolved molecular structure | ~1 mg | Solid (or liquid droplet trapped on transparent substrate) |  | Mount on IR-transparent window (e.g. KBr ZnSe). ZnSe typically used for ATR | Clean sample surface; avoid overlapping contaminants | ng–µg | IR microscope with MCT (Mercury Cadmium Telluride) or DTGS detector | Contaminant identification microplastics polymer inclusions multilayer films | Enables IR analysis of microscopic regions; often used with FTIR mapping or imaging; ATR or transmission modes available depending on setup
Raman Spectroscopy | Raman | Small molecules Polymers | Functional group analysis | 1-100 mg | Solid or liquid | Water Methanol Ethanol Acetonitrile Chloroform Toluene | Minimal preparation; clean surface for solids | Fluorescence-free | μg-mg | Laser (e.g. 532 nm 785 nm) with CCD detector | Purity ID | Non-destructive; complementary to FTIR; less sensitive to water than FTIR
Karl Fischer Titration | KF | Small molecules | Water content determination | 10-100 mg | Solid or liquid | Compatible with KF reagents (e.g. methanol-based) | Dissolve or disperse in KF reagent | Compatible with reagents | ppm | Potentiometric or coulometric titration | QC pharmaceuticals | Highly specific for water content; volumetric or coulometric methods available
Inductively Coupled Plasma Mass Spectrometry | ICP-MS | Small molecules Polymers | Elemental analysis | 1-10 mg | Solution | Aqueous or dilute acid (e.g. nitric acid) | Digest in acid (e.g. HNO₃) filter to remove particulates | Free from particulates | pg-ng | Quadrupole or sector field MS | QC elemental analysis | Trace and ultra-trace elemental analysis; requires complete sample digestion
Differential Scanning Calorimetry | DSC | Small molecules Polymers | Thermodynamic analysis (glass transition melting crystallization curing) | 0.5-10 mg | Solid | N/A | Place dry homogeneous sample in aluminum or sealed crucible | Representative sample | μg–mg | Heat flow sensors (differential thermocouples) | QC research Tg Tm Tc crosslinking behavior purity assessment | Baseline correction and sample sealing are crucial; often coupled with TGA or DMTA for full thermal-mechanical profiling; especially useful for polymer and pharmaceutical analysis
Surface Plasmon Resonance | SPR | Biomolecules Small molecules | Interaction analysis | ~0.01–1 mg or 10–100 µL (solution) | Solution | Aqueous buffers (e.g. PBS HEPES Tris) or compatible solvents for small molecules | Filter to remove aggregates; dilute to appropriate concentration | High purity free from aggregates and contaminants | ng-µg (dependent on binding affinity and sensor chip) | SPR sensor chip flow system refractive index measurement | Drug discovery biochemistry | Measures real-time binding kinetics and affinity; requires stable sensor chip surface
Supercritical Fluid Chromatography Mass Spectrometry | SFC-MS | small molecules chiral compounds lipids | Chiral/non-polar separation & ID | 0.01-1 mg | Compatible with supercritical fluid | Methanol ethanol isopropanol supercritical CO₂ compatible solvents | Dissolve in suitable solvent filter degas | High purity | pg-ng | MS (APCI or ESI) UV | Pharmaceuticals research | 1. Offers faster and greener separations than LC 2. Suitable for non-polar to moderately polar analytes
Time-of-Flight Mass Spectrometry | TOF-MS | Small molecules polymers | High-speed high-resolution mass analysis | 0.01-1 mg | Soluble in suitable solvent (water acetonitrile methanol) |  | Dissolve in suitable solvent filter remove salts | Free from salts buffers detergents | Sensitivity depends on instrument configuration and ionization method typically pg-ng | Time-of-flight mass analyzer | Research metabolomics polymer mass determination | 1. Offers high-resolution fast acquisition 2. Ideal for complex mixtures and exact mass analysis
Size Exclusion Chromatography | SEC or GPC | Polymers | Determination of molecular weight distribution | ~1–10 mg | Solution in compatible eluent (e.g. THF DMF water) | Depends on column chemistry and polymer; THF DMF water commonly used | Filter to remove particles; dissolve fully in eluent | free of aggregated and insoluble material | μg-mg | RI UV MALS (multi-angle light scattering) | Polymer characterization protein size distribution nanoparticles | 1. Separation based on hydrodynamic volume not molecular weight directly 2. Requires calibration for MW analysis
X-ray Diffraction | XRD | Small molecules polymer thin films | Phase identification crystallinity structure determination | ~1–50 mg powder or thin solid layer | Solid (Powder or thin film) | N/A | Fine powder pressed onto holder; flat thin film adhered to substrate | Free from moisture and surface contamination | 1–5% wt% | X-ray detector (scintillation or solid-state) | Materials identification crystal structure polymer crystallinity | 1. Only suitable for crystalline materials 2. Amorphous phases are undetectable unless coupled with other techniques
Thermal Gravimetric Analysis | TGA | Small molecules Polymers | Thermal stability decomposition profile compositional analysis | ~5–20 mg | Solid (Powder film pellet) | N/A | Dry and weigh the sample; place in crucible | Free from moisture and volatile contaminants | μg-mg | Microbalance + Furnace | Polymer degradation filler content moisture content analysis | Can be coupled with MS/FTIR for evolved gas analysis (TGA-MS or TGA-FTIR)
Scanning Electron Microscopy | SEM | Small molecules Polymers | Surface morphology and topography analysis | ~1–10 mg size typically <1 cm² | Solid | N/A | Conductive coating (e.g. gold platinum) if non-conductive; mounted on stub | Clean dry surface free of contaminants | Nanometer resolution (surface structure) | Secondary electron detector backscattered electron detector | Polymers nanomaterials surfaces coatings fracture analysis | 1. Requires vacuum 2. Charging may occur for non-conductive samples without coating
Energy Dispersive X-ray Spectroscopy (often coupled with SEM) | EDX or EDS | Small molecules Polymers | Elemental composition analysis (qualitative and semi-quantitative) | 1–10 mg | Solid (bulk film or powder mounted on stub) | N/A | Dry clean flat surface; conductive coating may be required if non-conductive | Free from surface contamination and moisture | ~0.1–1 wt% varies by element and instrument | Energy-dispersive X-ray detector (usually attached to SEM) | Elemental mapping failure analysis materials characterization | Detects elements from ~Z=5 (boron) and above; often used with SEM for compositional contrast
Dynamic Mechanical Thermal Analysis | DMTA or DMA | Polymers | Measurement of mechanical properties as a function of temperature frequency or time | ~10–50 mg | Solid (molded or cut sample) | N/A | Prepare standardized shape and dimensions; dry if hygroscopic | Free of surface contamination and voids | µN-N force; nm scale displacement | Displacement sensors + temperature control system | Glass transition viscoelastic properties (modulus tan δ) damping | Offers insight into mechanical behavior under realistic stress/strain and temperature; requires precise geometry and clamping
Differential Thermal Analysis | DTA | Small molecules Polymers | Thermal transitions: melting crystallization decomposition | ~5–20 mg | Solid (powder film or pellet) | N/A | Dry and weigh sample; load into crucible alongside reference material | Free of moisture and volatile impurities | μg–mg | Thermocouples measuring temperature difference between sample and reference | Melting point determination phase transitions thermal stability | Often performed alongside TGA (simultaneous TG-DTA); does not measure heat flow only temperature differences
Transmission Electron Microscope | TEM | Polymers | Internal structure imaging at atomic/nanometer scale | ~1 mg | Solid (ultra-thin film or embedded and microtomed sample) | N/A | Ultramicrotomy focused ion beam (FIB) or drop-casting on grid; must be vacuum-stable | Free from dust oils and aggregation; high purity needed for clarity | Sub-nanometer resolution | CCD/CMOS cameras electron diffraction detectors | Crystal structure grain boundaries nanoparticle morphology defect analysis | Requires extremely thin samples; sensitive to beam damage; imaging under high vacuum and electron beam acceleration (80–300 keV)
Atomic Force Microscope | AFM | Polymers thin films | Surface topography roughness mechanical and electrical surface properties | ~1 mg | Solid (film crystal substrate-coated material) | N/A | Mount on magnetic disk or stub; clean flat surface essential | Free from dust loose particles and fingerprints | Sub-nanometer vertical resolution | Laser deflection system with photodiode | Surface roughness phase imaging stiffness mapping nanostructures | Can operate in air liquid or vacuum; multiple modes (tapping contact non-contact force spectroscopy) allow broad versatility
Nanoindentation |  | Small molecules Polymers | Measurement of mechanical properties: hardness elastic modulus creep | ~1 mg | Solid (flat films bulk materials coatings) | N/A | Polish if needed; ensure flatness cleanliness and proper adhesion (for films) | Free of dust and surface contamination | nN-µN force; nm depth resolution | Load-displacement sensors + piezo stage | Thin film hardness/modulus viscoelasticity time-dependent deformation | Requires precise surface preparation; results sensitive to substrate effect in thin films (usually <10% rule for indentation depth)
Electron Paramagnetic Resonance/Electron Spin Resonance | EPR or ESR | Small molecules Polymers | Detection and characterization of paramagnetic species | ~10–50 µL or ~5–20 mg | Solid or solution | Depends on sample; solvents must be EPR-compatible (low dielectric loss) | Degas if needed; load into quartz EPR tubes (~4 mm outer diameter) | Avoid diamagnetic contaminants; remove oxygen for sensitive systems | ~10⁻⁹ to 10⁻¹² mol | Microwave bridge and magnetic field sweep | Radical detection transition metal analysis redox chemistry photophysics | Highly selective for unpaired electrons; complementary to NMR for metal centers and radicals; samples often measured at low temperature (e.g. liquid nitrogen or helium cooled)
Small-Angle X-ray Scattering | SAXS | Small molecules Polymers | Structural analysis at the nanoscale (1–100 nm); particle size shape and distribution | ~0.5–1 mL | Liquid (solution) powder or solid film | For solution SAXS: buffer or solvent must be compatible with X-ray beam and capillaries (e.g. water ethanol) | Filter and degas liquid samples; fill into quartz capillary or flow cell | No dust or large aggregates; matched solvent background | Sensitive to differences in electron density (~0.1% contrast) | 2D X-ray detector (CCD or hybrid pixel detectors) | Nanoparticle size distribution micelle structure polymer domain spacing protein conformation | Non-destructive; provides averaged structural data; often performed at synchrotron for high resolution and time-resolved studies
Synchrotron Radiation | SR (e.g. SXRD SAXS) | Small molecules Polymers Nanomaterials | High-resolution structural analysis diffraction scattering | ~0.1–10 mg or 0.1–1 mL (solution powder or thin film); for solution SAXS use quartz capillaries; for SXRD powders or films | Solution powder or thin film | For solution SAXS: buffer or solvent must be compatible with X-ray beam and capillaries (e.g. water ethanol) | Filter and degas for solution SAXS; powders pressed or films adhered for SXRD | No dust or large aggregates; matched solvent background for solution SAXS | Sensitive to differences in electron density (~0.1% contrast) | Synchrotron X-ray source 2D detectors (e.g. CCD pixel detectors); energy range 5–100 keV | Crystallography nanoparticle characterization polymer morphology time-resolved studies | Requires access to synchrotron facilities; high brilliance enables fast high-resolution data collection
FIB cut | FIB | Small molecules Polymers Materials | Sample preparation for SEM/TEM cross-sectional analysis | Solid (bulk film or device); size <1 cm²; vacuum-stable | Solid | N/A | Mount on stub; ensure vacuum compatibility | Free from dust and surface contamination | Nanometer-scale resolution for milling precision | Focused ion beam (e.g. Ga⁺ ions) 5–30 keV beam energy | Microelectronics materials failure analysis TEM sample preparation | Destructive; precise milling for nanoscale features; often coupled with SEM (FIB-SEM)
BET | BET | Small molecules Polymers Powders | Surface area and porosity analysis | ~10–100 mg (powder or porous solid); dry degassed | Powder or porous solid | N/A | Degas at elevated temperature under vacuum | Free of adsorbed gases and contaminants | Sensitive to surface area changes (~0.1 m²/g) | Nitrogen adsorption at 77 K; BET model for surface area calculation | Catalyst characterization nanomaterial analysis polymer porosity | Requires degassing to remove adsorbed gases; sensitive to sample preparation